## Slide 1: Measurement noise

## Slide 2: Measurement Noise
Noise is added always over the measurements.
The extent to which series mode noise corrupts measurement signals is measured by a quantity known as the
signal-to-noise ratio. This is defined as:
Signal-to-noise ratio = 20 log (Vs / Vn ) where
Vs is the mean voltage level of the signal and
Vn is the mean voltage level of the noise.
In the case of a.c. noise voltages, the root-mean squared value is used as the mean.

## Slide 3: Sources of measurement noise
Inductive coupling
Capacitive (electrostatic) coupling
Noise due to multiple earths
Noise in the form of voltage transients
Thermoelectric potentials
Shot noise
Electrochemical potentials

## Slide 4: Sources of measurement noise
Inductive coupling
If signal-carrying cables are close to such external cables or equipment, a significant mutual inductance M can exist between them, and this can generate a series mode noise voltage of several millivolts given by Vn = M I, where I is the rate of change of current in the mains circuit.

## Slide 5: Sources of measurement noise
Capacitive (electrostatic) coupling
Capacitive coupling, also known as electrostatic coupling, can also occur between the signal wires in a measurement circuit and a nearby mains-carrying conductor. The series mode noise voltage Vn is zero if the coupling capacitances are perfectly balanced, i.e. if C1 = C2 and C3 = C4. However, exact balance is unlikely in practice,

## Slide 6: Sources of measurement noise
Noise due to multiple earths
Measurement signal circuits are isolated from earth.
However, leakage paths often exist between measurement circuit signal wires and earth at both the source (sensor) end of the circuit and also the load (measuring instrument) end.
This does not cause a problem as long as the earth potential at both ends is the same.
However, it is common to find that other machinery and equipment carrying large currents is connected to the same earth plane. This can cause the potential to vary between different points on the earth plane. This situation, which is known as multiple earths, can cause a series mode noise voltage in the measurement circuit.

## Slide 7: Sources of measurement noise
Noise in the form of voltage transients
When motors and other electrical equipment (both a.c. and d.c.) are switched on and off, large changes of power consumption suddenly occur in the electricity supply system.
This can cause voltage transients (‘spikes’) in measurement circuits connected to the same power supply. Such noise voltages are of large magnitude but short time duration.
Corona discharge can also cause voltage transients on the mains power supply. This occurs when the air in the vicinity of high voltage d.c. circuits becomes ionized and discharges to earth at random times.

## Slide 8: Sources of measurement noise
Thermoelectric potentials
Whenever metals of two different types are connected together, a thermoelectric potential (sometimes called a thermal e.m.f.) is generated according to the temperature of the joint.
This is known as the thermoelectric effect and is the physical principle on which temperature-measuring thermocouples operate.
Such thermoelectric potentials are only a few millivolts in magnitude and so the effect is only significant when typical voltage output signals of a measurement system are of a similar low magnitude.

## Slide 9: Sources of measurement noise
Shot noise
Shot noise occurs in transistors, integrated circuits and other
semiconductor devices.
It consists of random fluctuations in the rate of transfer of carriers across junctions within such devices.

## Slide 10: Sources of measurement noise
Electrochemical potentials
These are potentials that arise within measurement systems due to electrochemical action.
Poorly soldered joints are a common source.

## Slide 11: Techniques for reducing measurement noise
Location and design of signal wires
Earthing
Shielding
Other techniques

## Slide 12: Techniques for reducing measurement noise
Location and design of signal wires
Both the mutual inductance and capacitance between signal wires and other cables are inversely proportional to the square of the distance between the wires and the cable.
Thus, noise due to inductive and capacitive coupling can be minimized by ensuring that signal wires are positioned as far away as possible from such noise sources.
A minimum separation of 0.3m is
essential, and a separation of at least 1m is preferable.
Noise due to inductive coupling is also substantially reduced if each pair of signal wires is twisted together.

## Slide 13: Techniques for reducing measurement noise
Earthing
Noise due to multiple earths can be avoided by good earthing practices. In particular, this means keeping earths for signal wires and earths for high-current equipment entirely separate. Recommended practice is to install four completely isolated earth circuits as follows:
Power earth: provides a path for fault currents due to power faults.
Logic earth: provides a common line for all logic circuit potentials.
Analogue earth (ground): provides a common reference for all analogue signals.
Safety earth: connected to all metal parts of equipment to protect personnel should power lines come into contact with metal enclosures.

## Slide 14: Techniques for reducing measurement noise
Shielding
Shielding consists of enclosing the signal wires in an earthed, metal shield that is itself isolated electrically from the signal wires.
The shield should be earthed at only one point, preferably the signal source end. A shield consisting of braided metal eliminates 85% of noise due to capacitive coupling whilst a lapped metal foil shield eliminates noise almost entirely.
The wires inside such a shield are normally formed as a twisted pair so that protection is also provided against induced noise due to nearby electromagnetic fields.
Metal conduit is also sometimes used to provide shielding from capacitve-coupled noise, but the necessary supports for the conduit provide multiple earth points.

## Slide 15: Techniques for reducing measurement noise
Other Methods
The phase-locked loop is often used as a signal-processing
element to clean up poor quality signals.
Lock-in amplifiers are also commonly used to extract d.c. or
slowly varying measurement signals from noise.

## Slide 16: Signal processing
Signal processing is concerned with improving the quality of the reading or signal at the output of a measurement system, and one particular aim is to attenuate any noise in the measurement signal that has not been eliminated by careful design of the measurement system
Analog signal processing
Digital signal processing

## Slide 17: Signal processing
Analog Signal Processing
Signal Integration
Voltage follower (pre-amplifier)
Voltage comparator
Phase-sensitive detector
Lock-in amplifier
Signal addition
Signal multiplication
Filters
  Passive
  Active
Signal amplification
Signal attenuation
Differential amplification
Signal linearization
Bias (zero drift) removal

## Slide 18: Signal processing
Digital signal processing
Signal Integration
Sample and hold circuit
Analog-to-digital (A/D) conversion
Digital-to-analog (D/A) conversion
Digital filtering
Autocorrelation
Other digital signal processing operations